## digitalSensor_addressMap
no | REGISTER | ADDRESS | REGISTER NAME | properties | Data Type | Data Length | 비고 | Sensor Command | Modbus Function code
1 | 40001 | 0 | Sensor Value | READ | FLOAT | 2 | 32bit IEEE-754 Floating Point | GSNSR | 0x03 : Read Holding Registers
2 | 40003 | 2 | Temperature Value | READ | FLOAT | 2 |  | GTEMP | 0x03 : Read Holding Registers
3 | 40005 | 4 | Sensor mV | READ | FLOAT | 2 |  | GSNSRMV | 0x03 : Read Holding Registers
4 | 40007 | 6 | Sensor Type | READ | UINT16 | 1 | 0 = 센서 없음/정의되지 않음 1 = pH 센서 2 = ORP 센서 3 = DO(용존산소) 센서 5 = 비접촉식 전도도 센서(Non-Contacting Conductivity) 6 = 오존 센서 7 = 저범위 탁도 센서(Low-Range Turbidity) 8 = 고범위 탁도 센서(High-Range Turbidity, Argus color, 사용되지 않음) 9 = 고범위 탁도 센서(High-Range Turbidity, Argus IR) 10 = 염소 센서 11 = 부유 고형물 센서(Suspended Solids, Argus IR) | GSTYPE | 0x03 : Read Holding Registers
5 | 40008 | 7 | Calibration Status | READ | UINT16 | 1 | 0 = No Sensor Calibration 1 = Cal in progress 2 = Cal OK 3 = Fal Not Stable 4 = Fail Buffer not found 5 = Fail 1st Buffer not found 6= Fail 2nd buffer not found 7= Value too low 8= Fail Value too high 9= Fali Slope too low 10= Fali Slope too high | CALSTATUS | 0x03 : Read Holding Registers
6 | 40009 | 8 | Sensor Serial Number GET | READ | UINT16 | 3 | 6byte 아스키 코드 | GSSNO | 0x03 : Read Holding Registers
7 | 40012 | 11 | Sensor Factor | R/W | FLOAT | 2 | adjust_value = slope * real_value + offset | 없음 | 0x03 : Read Holding Registers 16 (0x10): Write Multiple Registers
8 | 40014 | 13 | Sensor Offset | R/W | FLOAT | 2 |  | 없음 | 0x03 : Read Holding Registers 16 (0x10): Write Multiple Registers
10 | 40016 | 15 | 1 Point Sample Calibration | R/W | FLOAT | 2 | 32bit IEEE-754 Floating Point | CALS1PB | 0x03 : Read Holding Registers 16 (0x10): Write Multiple Registers
11 | 40018 | 17 | 2 Point Buffer Calibration | write | UINT16 | 1 | first : 0, second : 1 | CALS2PB | 0x03 : Read Holding Registers 06 (0x06): Write Single Register
5 | 40019 | 18 | Zero Calibration | write | UINT16 | 1 | Returns nothing | CALZERO | 0x03 : Read Holding Registers 06 (0x06): Write Single Register
6 | 40020 | 19 | Calibration Abort | write | UINT16 | 1 | Returns nothing | CALABORT | 0x03 : Read Holding Registers 06 (0x06): Write Single Register
float 범위
1. 양의 범위 : 0x00800000 ~ 0x7F7FFFFF
2. 음의 범위 : 0xFF7FFFFF ~ 0x80800000
3. 특수값 : 0: 0x00000000 -0: 0x80000000
4. 무한대(Infinity): 0x7F800000
5. 음의 무한대(-Infinity): 0xFF800000
6. NaN(Not a Number): 0x7FC00000
UINT16 범위
0 ~ 65,535 (0x0000 ~ 0xFFFF)