## digitalSensor_addressMap
no | REGISTER | ADDRESS | REGISTER NAME | properties | Data Type | data Length (1은 16bit ) | 비고 | Modbus Function code
1 | 40007 | 6(0x06) | Sensor Value | READ | FLOAT | 2 | 32bit IEEE-754 Floating Point | 0x03 : Read Holding Registers
2 | 40009 | 8(0x08) | Temperature Value | READ | FLOAT | 2 |  | 0x03 : Read Holding Registers
3 | 40017 | 16(0x10) | Sensor mV | READ | FLOAT | 2 |  | 0x03 : Read Holding Registers
4 | 40048 | 47(0x2f) | Sensor Type | READ | UINT16 | 1 | 0 = 센서 없음/정의되지 않음 1 = pH 센서 2 = ORP 센서 3 = DO(용존산소) 센서 4 = 접촉식 전도도 센서 5 = 비접촉식 전도도 센서(Non-Contacting Conductivity) 6 = 오존 센서 7 = 저범위 탁도 센서(Low-Range Turbidity) 8 = 고범위 탁도 센서(High-Range Turbidity, Argus color, 사용되지 않음) 9 = 고범위 탁도 센서(High-Range Turbidity, Argus IR) 10 = 염소 센서 11 = 부유 고형물 센서(Suspended Solids, Argus IR) | 0x03 : Read Holding Registers
5 | 40002 | 1(0x01) | Calibration Status | READ | UINT16 | 1 | 0 or 0x0e = No Sensor Calibration 1 = Cal in progress 2 = Cal OK 3 = Fal Not Stable 4 = Fail Buffer not found 5 = Fail 1st Buffer not found 6= Fail 2nd buffer not found 7= Value too low 8= Fail Value too high 9= Fali Slope too low 10= Fali Slope too high | 0x03 : Read Holding Registers
6 | 40050 | 49(0x31) | Sensor Serial Number GET | READ | UINT16 | 3 | 6byte 아스키 코드 | 0x03 : Read Holding Registers
7 | 40084 | 83(0x53) | Sensor Factor | R/W | FLOAT | 2 | 32bit IEEE-754 Floating Point | 0x03 : Read Holding Registers 16 (0x10): Write Multiple Registers
8 | 40086 | 85(0x55) | Sensor Offset | R/W | FLOAT | 2 |  | 0x03 : Read Holding Registers 16 (0x10): Write Multiple Registers
10 | 40011 | 10(0x0a) | 1 Point Sample Calibration | R/W | FLOAT | 2 | 32bit IEEE-754 Floating Point | 16 (0x10): Write Multiple Registers
11 | 40013 | 12(0x0c) | 2 Point Buffer Calibration | write | UINT16 | 1 | first : 0, second : 1 | 06 (0x06): Write Single Register
5 | 40023 | 22(0x16) | Zero Calibration | write | UINT16 | 1 | Returns nothing | 06 (0x06): Write Single Register
6 | 40024 | 23(0x17) | Calibration Abort | write | UINT16 | 1 | Returns nothing | 06 (0x06): Write Single Register
7 | 40026 | 25(0x19) | Change Slave ID | R/W | UINT16 | 1 | Master에서 Slave ID를 "250"으로 국번 변경 및 읽을 수 있음 | 0x03 : Read Holding Registers 06 (0x06): Write Single Register
float 범위
1. 양의 범위 : 0x00800000 ~ 0x7F7FFFFF
2. 음의 범위 : 0xFF7FFFFF ~ 0x80800000
3. 특수값 : 0: 0x00000000 -0: 0x80000000
4. 무한대(Infinity): 0x7F800000
5. 음의 무한대(-Infinity): 0xFF800000
6. NaN(Not a Number): 0x7FC00000
UINT16 범위
0 ~ 65,535 (0x0000 ~ 0xFFFF)